## Slide 1: The science behind your device
Jill Newton
Spruce Environmental Technologies
2019 International Radon Symposium Denver, Colorado

## Slide 2
All listed devices chamber-tested
Must meet requirements in MS-PC
Device Listing
2019 International Radon Symposium Denver, Colorado

## Slide 3
Good precision
Poor bias
Radon Level
Good precision
Good bias
X
Poor precision
Good bias
Accuracy = Precision and Bias
Good precision is like having a good “group” of shots. That is, every-thing is consistent…but it can be off target.
Good bias is when the average of group is close to bullseye…but some hits may be off the mark
Poor precision
Poor bias
2019 International Radon Symposium Denver, Colorado

## Slide 4
Sensitivity
How many of the radon decay events are actually counted (and not other events)
The more data your device can collect, the more certain you can be it’s giving an answer you can use to make a mitigation decision
Device Sensor

## Slide 5
Radon and decay products radioactively disintegrate
Emit radiation from nucleus
  Alpha α
  Beta β
  Gamma γ
Radon test devices count what these occurrences do
Radioactive Decay Reminder
2019 International Radon Symposium Denver, Colorado

## Slide 6
Radon Decay Reminder
Radon gas
γ
α
Solid decay products








2019 International Radon Symposium Denver, Colorado

## Slide 7
Charcoal Canisters
Liquid Scintillation
Alpha Track
Continuous Monitors
Electret Ion Chambers
Common Radon Devices

## Slide 8
Electret Ion Chamber
2019 International Radon Symposium Denver, Colorado

## Slide 9
Reading Electrets
+ + + + + +
Specialized Teflon® plate electrostatically charged positive
Multiple configurations
2019 International Radon Symposium Denver, Colorado

## Slide 10
Calculate Radon Reading
2019 International Radon Symposium Denver, Colorado

## Slide 11
γ
Why Gamma Subtraction?
2019 International Radon Symposium Denver, Colorado

## Slide 12
Dentist’s, doctor’s offices
Lab
Higher elevation
Sensitive to temperature & moisture
Leave in controlled setting
γ
Device-Specific Cautions
2019 International Radon Symposium Denver, Colorado

## Slide 13
Dust and fibers
Touch surface
Device-Specific Cautions
2019 International Radon Symposium Denver, Colorado

## Slide 14
10% duplicates
5% blanks
3% spikes
QA/QC Requirements
2019 International Radon Symposium Denver, Colorado

## Slide 15
Annual Calibration
Reference Electrets
Routine instrument checks
What does the manufacturer tell you to do?
QA/QC (Continued)
2019 International Radon Symposium Denver, Colorado

## Slide 16
Diffusion barrier
Open face
Test kit
Charcoal Devices
2019 International Radon Symposium Denver, Colorado

## Slide 17
Structure of Growing Woody Materials
Charcoal Structure
Water
Structure of growing woody materials
Charcoal Structure
The Nature of Charcoal
2019 International Radon Symposium Denver, Colorado

## Slide 18
Treated with heat and pressurized steam, oxygen, or other gas or with acids, bases or salts
Further opens pores gives greater surface area
One gram of charcoal
=
Area of 1-3 basketball courts
Activating Charcoal
=
Weight of 1 Paperclip
2019 International Radon Symposium Denver, Colorado

## Slide 19
Radon diffuses into tiny charcoal pores
Some radon will decay inside pores
Decay products ADsorb (i.e., stick to) onto the carbon surface (decay products electrostatically charged and slightly chemically reactive)
Charcoal Sampling
2019 International Radon Symposium Denver, Colorado

## Slide 20
If water molecules block or fill the holes in the charcoal, the radon cannot diffuse into that space.
The more water vapor in the charcoal, the fewer holes there are for radon to get into and the fewer decay products that are left behind in the charcoal.
H2O
Humidity & Charcoal Caution
H2O
2019 International Radon Symposium Denver, Colorado

## Slide 21
Open face
2 - 3 days
Diffusion barrier
5 - 7 days
Device-Specific Cautions: Charcoal
2019 International Radon Symposium Denver, Colorado

## Slide 22
Radiator
Hot water heater
Electronic equipment
Sunshine
Heat & Charcoal
2019 International Radon Symposium Denver, Colorado

## Slide 23
Fan
Dehumidifier exhaust
HVAC vent
Air & Charcoal
2019 International Radon Symposium Denver, Colorado

## Slide 24
Because of diffusion, adsorption and desorption, charcoal canisters may be slightly biased to the last 12 to 24 hours of the test period.
Charcoal Bias
2019 International Radon Symposium Denver, Colorado

## Slide 25
Securely replace tape after deployment
Use only tape lab has supplied
Tape Cautions
2019 International Radon Symposium Denver, Colorado

## Slide 26
Return to lab immediately
Radioactive decay continues after device is retrieved
Shipping Charcoal Cautions
2019 International Radon Symposium Denver, Colorado

## Slide 27
Gamma radiation from decay products
Lab Analysis
2019 International Radon Symposium Denver, Colorado

## Slide 28
10% duplicates
5% blanks
3% spikes
Charcoal QA/QC Requirements
2019 International Radon Symposium Denver, Colorado

## Slide 29
Sample as charcoal
Cautions same as charcoal canisters
Testing period 2-4 days
Liquid Scintillation
2019 International Radon Symposium Denver, Colorado

## Slide 30
Liquid scintillant added to charcoal
Decay products DEsorb into liquid
Lab Analysis
2019 International Radon Symposium Denver, Colorado

## Slide 31
Radioactive decay causes liquid to emit light pulses
Photomultiplier detects pulses
Lab Analysis (Continued)
2019 International Radon Symposium Denver, Colorado

## Slide 32
10% duplicates
5% blanks
3% spikes
LS QA/QC
2019 International Radon Symposium Denver, Colorado

## Slide 33
Alpha Tracks
2019 International Radon Symposium Denver, Colorado

## Slide 34
Plastic coated strip
Coated Plastic
2019 International Radon Symposium Denver, Colorado

## Slide 35
Track Analysis Systems
Lab Analysis
2019 International Radon Symposium Denver, Colorado

## Slide 36
Plastic coating pocked by alpha particles
Slide soaked in caustic solution to enlarge damage marks on film
Tracks of Alpha Particles
2019 International Radon Symposium Denver, Colorado

## Slide 37
5% blanks
3% spikes
Alpha Track QA/QC
10% duplicates
2019 International Radon Symposium Denver, Colorado

## Slide 38
Long term
91 days to 1 year
Device- Specific Cautions
2019 International Radon Symposium Denver, Colorado

## Slide 39
Their End
  Rigorous QC
  Licensing & affiliations
  User friendly interfaces
  Timely turnaround
  Responsiveness
  Help
Your End
  Accurate chain of custody
  Complete chain of custody
  Timely and complete QC
  Following instructions
  Thorough record-keeping
You and Your Lab
2019 International Radon Symposium Denver, Colorado

## Slide 40
Chain of Custody

## Slide 41
Chain of Custody

## Slide 42
Chain of Custody

## Slide 43
WPB Radon Testing Bethlehem, PA
Alpha scintillation
Pulsed ion
Solid state silicon
Continuous Radon Monitors
2019 International Radon Symposium Denver, Colorado

## Slide 44
Alpha scintillation
Cell lined with zinc sulfide
Zinc Sulfide- Lined Chamber
WPB Radon Testing Bethlehem, PA
2019 International Radon Symposium Denver, Colorado

## Slide 45
Pulses of Light
Zinc Sulfide
Radon nucleus
Scintillation cell
2019 International Radon Symposium Denver, Colorado

## Slide 46
Negative
Positive
+
−
Electrometer counter
Continuous Radon Monitor: Pulsed Ion Chamber
2019 International Radon Symposium Denver, Colorado

## Slide 47
Silicon
Chip
Electrical
Pulse Counter
Continuous Radon Monitor: Solid State Silicon Detector
2019 International Radon Symposium Denver, Colorado

## Slide 48
Temperature
Humidity
Barometric pressure
Tamper (movement)
AC Interrupt
Bells & Whistles
2019 International Radon Symposium Denver, Colorado

## Slide 49
Temperature
Humidity
Barometric Pressure
Motion
Hours
Radon Level
Femtotech
Getting Results
2019 International Radon Symposium Denver, Colorado

## Slide 50
Results
2019 International Radon Symposium Denver, Colorado

## Slide 51
Self Check 10% Duplicates Annual Calibration
Cross Checks
CRM QA/QC
2019 International Radon Symposium Denver, Colorado

## Slide 52
Homeowner Devices
2019 International Radon Symposium Denver, Colorado

## Slide 53
Measurement time period
Humidity, temperature, air movement
Sensitivity
Cost
Analysis – background subtractions & correction factors
Strengths & Cautions
2019 International Radon Symposium Denver, Colorado

## Slide 54
What’s in Your Files?
Quality control measures
  Duplicates
  Field blanks
  Spikes
  Source checks
  Cross checks
  Background checks
  Routine instrument checks
  Calibration
Calculations
Control charts
2019 International Radon Symposium Denver, Colorado

## Slide 55: Questions?
Jill Newton
jnewton@spruce.com
2019 International Radon Symposium Denver, Colorado